VIDYUTHLABS ¦

PRODUCT DESIGN DOCUMENT

VL-STAT-v1 Potentiostat / EIS Analyzer

| Document Number | VL-HW-001 |
| --- | --- |
| Revision | 1.0 |
| Date | 08 April 2026 |
| Classification | CONFIDENTIAL — Internal Use Only |
| Author | Varshini CB |
| Reviewed By | VidyuthLabs Engineering Lead |
| Project Phase | PDR → CDR → Prototype Build |
| Target Fab | JLCPCB — 4-Layer ENIG |

This document contains proprietary information belonging to VidyuthLabs. Unauthorized reproduction or distribution is strictly prohibited.

© 2026 VidyuthLabs. All Rights Reserved.

Vidyuthlabs.co.in ¦ Bengaluru, India

# Document Control

## Revision History

| Rev | Date | Author | Description | Status |
| --- | --- | --- | --- | --- |
| 1.0 | 08 Apr 2026 | Varshini CB | Initial release — PDR/CDR combined document | Released |
| — | — | — | Reserved for v2 updates | Draft |

## Document Scope

This Product Design Document (PDD) covers the complete lifecycle from product requirements through hardware architecture, preliminary design, and critical design for the VidyuthLabs VL-STAT-v1 potentiostat and electrochemical impedance spectroscopy (EIS) analyzer. It is structured to follow standard hardware development phases: PRD → PDR → CDR → DFM, and serves as the single authoritative reference for all hardware design decisions.

## Definitions and Abbreviations

| Term | Definition |
| --- | --- |
| EIS | Electrochemical Impedance Spectroscopy |
| SPE | Screen-Printed Electrode (here: PCB-integrated) |
| TIA | Trans-Impedance Amplifier |
| WE | Working Electrode |
| CE | Counter Electrode |
| RE | Reference Electrode |
| AFE | Analog Front-End |
| LGA | Land Grid Array (IC package) |
| ENIG | Electroless Nickel Immersion Gold (PCB finish) |
| PDR | Preliminary Design Review |
| CDR | Critical Design Review |
| DFM | Design for Manufacturing |
| SPI | Serial Peripheral Interface |
| UART | Universal Asynchronous Receiver-Transmitter |
| RCAL | Resistor Calibration (AD5941 feature) |
| FR4 | Standard PCB laminate substrate material |

# Section 1 — Product Requirements Document (PRD)

| PRD | This section defines what the product must do, for whom, and under what constraints. All subsequent design decisions must be traceable to requirements defined here. |
| --- | --- |

## 1.1  Product Vision

VL-STAT-v1 is a compact, USB-C powered, wirelessly connected potentiostat and EIS analyzer designed for research and diagnostic applications in biosensing, corrosion analysis, battery characterization, and environmental monitoring. It integrates a PCB-native three-electrode electrochemical cell (SPE) onto the same PCB as the measurement electronics, reducing system complexity, connector variability, and cost.

## 1.2  Stakeholder Analysis

| Stakeholder | Needs | Priority |
| --- | --- | --- |
| Research Scientist | High-accuracy impedance data; flexible excitation frequency range; Python/MATLAB export | HIGH |
| Field Engineer | Wireless connectivity; battery/USB power; compact form; robust enclosure | MEDIUM |
| Embedded Developer | Open firmware; SWD debug access; UART console; documented API | MEDIUM |
| Manufacturing | JLCPCB-compatible BOM; ENIG; ≤4 layer; standard DFM rules | HIGH |

## 1.3  Functional Requirements

- FR-001: The system shall perform cyclic voltammetry (CV) with a scan rate of 1 mV/s to 1 V/s and potential range of ±1.5V.

- FR-002: The system shall perform EIS over a frequency range of 0.1 Hz to 200 kHz with a minimum of 7 frequency points per decade.

- FR-003: The system shall measure electrochemical current with a minimum resolution of 10 nA using the AD5941 on-chip TIA.

- FR-004: The system shall stream measurement data wirelessly via IEEE 802.11 b/g/n (Wi-Fi) and BLE 5.0 simultaneously.

- FR-005: The system shall be powered from USB-C 5V/500mA without external power supplies.

- FR-006: The PCB shall integrate a three-electrode SPE (WE, CE, RE) with an ENIG-finished gold working electrode.

- FR-007: The microcontroller shall support SWD/JTAG debugging and field firmware updates over USB.

- FR-008: The system shall execute RCAL self-calibration on power-up and on temperature change exceeding 5°C.

## 1.4  Non-Functional Requirements

| ID | Category | Requirement | Target |
| --- | --- | --- | --- |
| NFR-001 | Performance | EIS impedance accuracy | ±2% magnitude, ±1° phase |
| NFR-002 | Performance | Measurement noise floor | < 1 nA RMS at TIA input |
| NFR-003 | Physical | PCB dimensions | 80 × 60 mm maximum |
| NFR-004 | Physical | Board stackup | 4-layer, 1.6 mm thickness |
| NFR-005 | EMC | Wi-Fi isolation from AFE | 15 mm copper-free RF keepout |
| NFR-006 | Manufacturing | PCB finish | ENIG on all layers |
| NFR-007 | Operating Conditions | Operating temperature | 0°C to 50°C |
| NFR-008 | Reliability | MTBF target | > 10,000 hours @ 25°C |

## 1.5  Constraints and Assumptions

- CONST-001: The design must use the AD5941 AFE IC as the primary measurement engine. No alternative AFEs are in scope for v1.

- CONST-002: The RA4M1 (Arduino R4 Minima module) is used for v1 prototyping. Migration to bare RA4M1 chip is deferred to v2.

- CONST-003: All components must be JLCPCB-compatible or directly sourceable from Digi-Key/Mouser/LCSC.

- CONST-004: CE/FCC compliance is not in scope for v1; pre-compliance testing is planned for v3.

- ASSUME-001: Nanomaterial coating of the WE is performed post-fabrication in-house; adhesion to ENIG gold is assumed suitable.

# Section 2 — Preliminary Design Review (PDR)

| PDR | This section captures the top-level architecture decisions, component selection rationale, and block-level design choices made prior to detailed schematic capture. |
| --- | --- |

## 2.1  System Architecture Overview

The VL-STAT-v1 is built around a three-IC architecture: (1) the AD5941 electrochemical AFE handles all analog signal generation and measurement autonomously; (2) the Renesas RA4M1 microcontroller acts as the SPI master, data orchestrator, and USB interface; and (3) the ESP32-S3-MINI provides all wireless connectivity without interfering with measurement timing.

| Subsystem | Selected Component | Rationale |
| --- | --- | --- |
| Analog Front-End | AD5941BCPZ (Analog Devices) | Integrated DDS, TIA, 16-bit ADC; EIS-optimized sequencer; industry reference |
| Microcontroller | Renesas RA4M1 (via Arduino R4 Minima) | ARM Cortex-M4, HW SPI/UART, Arduino ecosystem for rapid prototyping |
| Wireless Module | ESP32-S3-MINI-1 (Espressif) | Dual-mode Wi-Fi/BLE 5.0; module-level FCC cert; minimal external BOM |
| Analog Power Reg | TLV1117-3.3 / LP2985-33 | Ultra-low noise (50 µV RMS); critical for EIS noise floor; LDO for low dropout |
| Digital Power Reg | AMS1117-3.3 | Cost-effective; noise less critical on digital rail |
| USB-C Connector | GCT USB4135-GF-A (or equiv.) | Standard USB-C receptacle; no USB PD required for v1 @ 5V/500mA |
| SPE Electrode | PCB-integrated, ENIG gold | Eliminates external connector variability; ENIG is validated WE surface for nanomaterial coating |

## 2.2  Inter-IC Communication Architecture

### 2.2.1  AD5941 ↔ RA4M1 (SPI)

The AD5941 is operated as an SPI peripheral with the RA4M1 as SPI master. The link runs as a 4-wire SPI interface at up to 16 MHz. The AD5941 INT pin (active-low) is connected to an external interrupt input on the RA4M1 (INTC0), enabling interrupt-driven data readout without polling. All SPI lines carry 33 Ω damping resistors at the RA4M1 end to suppress ringing on the PCB traces.

### 2.2.2  RA4M1 ↔ ESP32-S3 (UART)

Full-duplex UART at 921,600 baud is used between the RA4M1 and ESP32-S3. The RA4M1 encodes measurement data as COBS-framed binary packets for reliable transmission without start/stop byte ambiguity. The ESP32 handles all wireless stack operations independently, decoupling RF timing from measurement timing completely.

## 2.3  Power Architecture

The power system derives two independent 3.3V rails from the USB-C 5V input. A 500 mA resettable polyfuse and Schottky protection diode protect the VBUS line. Two separate LDOs provide AVDD (analog, ultra-low noise) and DVDD (digital). A 600Ω @ 100 MHz ferrite bead is inserted between the AVDD LDO output and the AD5941 AVDD pin to attenuate any digital switching noise coupling onto the analog supply.

The total power budget is: AD5941 ~15 mA + RA4M1 ~50 mA + ESP32-S3 peak ~240 mA (Wi-Fi TX) = ~305 mA, well within the 500 mA USB-C budget, providing a comfortable 195 mA margin.

## 2.4  PCB Architecture

The board is 80 × 60 mm with R4 corner radii, four M3 mounting holes, and a bottom-edge USB-C connector. A 4-layer stackup (Signal / GND / Power / Signal+GND) provides excellent analog isolation. The board is divided into three functional zones: Analog (top-left, AD5941 and power), Digital (bottom-left, RA4M1 and USB-C), RF (bottom-right corner, ESP32-S3 with 15 mm antenna keepout), and Electrode (right half, SPE pads with guarded traces).

## 2.5  PDR Open Items

| OI# | Open Item | Owner | Target Close |
| --- | --- | --- | --- |
| OI-001 | Confirm LP2985 availability at LCSC; verify stock before BOM lock | Varshini CB | PDR+2w |
| OI-002 | Characterize nanomaterial coating adhesion to ENIG vs OSP; run test coupons | Chemistry Team | PDR+4w |
| OI-003 | Confirm RA4M1 SPI timing compatibility with AD5941 at 16 MHz; prototype bench test | Varshini CB | CDR |

# Section 3 — Critical Design Review (CDR)

| CDR | This section captures detailed, implementation-ready hardware design decisions. All entries represent locked design choices ready for schematic capture and PCB layout. |
| --- | --- |

## 3.1  AD5941 Analog Front-End — Detailed Design

The AD5941BCPZ is housed in a 40-pin LGA (6 × 6 mm) package. The exposed thermal pad on the bottom of the LGA must connect to AGND via a 3 × 3 grid of 9 vias (0.3 mm drill) for thermal relief and ground reference quality. All AVDD bypass capacitors (10 µF + 100 nF per pin) must be placed within 0.5 mm of their respective IC pins on F.Cu; routing via inner layers defeats their purpose.

The RCAL calibration resistor (10 kΩ, 0.1% tolerance, 25 ppm/°C, Vishay TNPW series, 0402) must be placed within 5 mm of the RCAL0/1 pins on the same PCB layer. The RCAL routine is executed at startup and after every temperature delta > 5°C, correcting TIA gain error and ADC DC offset.

## 3.2  WE Node — High-Impedance Protection

The WE input (AD5941 pin 4, AIN0) operates in fA-level current measurement mode during low-current EIS. This is the most sensitive node on the board. The maximum permitted trace length from the WE pad to AD5941 AIN0 is 10 mm. The trace must be surrounded by a guard ring connected to the AD5941 GUARD output (pin 30), which drives the same potential as the WE input to eliminate surface leakage current.

| CRITICAL | No other signal trace may pass within 2 mm of the WE trace without a guard trace between them. This rule is MANDATORY and must be DRC-enforced in KiCad. |
| --- | --- |

## 3.3  SPE Electrode Geometry — Locked Dimensions

| Electrode | Geometry | Cu Layer | Mask Opening | Net Name |
| --- | --- | --- | --- | --- |
| Working (WE) | Ø 3.0 mm disk | F.Cu | Ø 3.12 mm | WE_PAD |
| Counter (CE) | Ring OD Ø7.0 / ID Ø5.0 mm | F.Cu | +0.12 mm both radii | CE_PAD |
| Reference (RE) | Ø 1.2 mm disk | F.Cu | Ø 1.32 mm | RE_PAD |
| WE–CE gap | 1.0 mm annular ring | No copper | N/A | N/A |
| RE offset from WE | 5.0 mm from WE center | — | — | — |

## 3.4  Power Subsystem — Decoupling Table

| Location | Value | Type | Placement Constraint |
| --- | --- | --- | --- |
| AD5941 AVDD (each pin) | 10 µF + 100 nF | X5R 0402/0603 | Within 0.5 mm of pin |
| AD5941 DVDD | 4.7 µF + 100 nF | X5R 0402 | Within 1 mm of pin |
| AD5941 VREF output | 10 µF | X5R 0603 | Within 1 mm of VREF pin |
| RA4M1 VDD | 100 nF/pin + 10 µF bulk | X5R/X7R | Per pin; bulk within 3 mm |
| ESP32-S3 VDD3P3 | 10 µF + 100 nF × 3 | Per Espressif ref design | Per ESP32-S3-MINI datasheet |
| LDO outputs | 10 µF each | 25V X5R ceramic | Within 1 mm of LDO output pin |

## 3.5  SPI and Trace Routing Rules

- Maximum SPI trace length: 50 mm for SCK, MOSI, MISO, CS. Length-matched to ±5 mm.

- 33 Ω series damping resistors (0402) on each SPI line, placed at the RA4M1 end.

- Minimum separation between SPI traces and WE/RE electrode traces: 3 mm. No parallel runs > 5 mm.

- UART traces (RA4M1 ↔ ESP32): route on B.Cu or through digital zone only. Max 100 mm, no constraints.

## 3.6  Grounding and Plane Strategy

A single, unbroken ground plane occupies L2 (In1.Cu). The plane is NOT split. Analog/digital current return path isolation is achieved through controlled component placement, not through physical plane cuts, which introduce inductance and can cause worse EMI than a unified plane.

AGND and DGND connect at a single star point located near the USB-C connector. The AD5941 AGND pins trace directly to this star node. Under no circumstances should digital return currents be permitted to flow beneath the AD5941 or the electrode zone.

## 3.7  Stackup Definition

| Layer | Name | Function | Notes |
| --- | --- | --- | --- |
| L1 | F.Cu | Signal + components + SPE copper | Primary routing; all SMD components |
| L2 | In1.Cu | Solid GND plane | Unbroken; 1 oz copper |
| L3 | In2.Cu | Power plane (AVDD / DVDD split) | Controlled split with star-point join |
| L4 | B.Cu | Signal + GND fill | Test points, UART traces, debug headers |

# Section 4 — Detailed Hardware Design

## 4.1  Power Block

The power input enters via the USB-C receptacle (GCT USB4135-GF-A or equivalent). Two CC pull-down resistors (5.1 kΩ to GND on each CC pin) signal to USB-C chargers that the device is a USB 2.0 sink requesting 5V/500 mA — this is mandatory per the USB-C specification and must not be omitted. A 500 mA resettable polyfuse (F1) and Schottky diode (D1) protect the VBUS rail before it reaches the two LDOs.

The AVDD LDO (U_REG1: TLV1117-3.3 or LP2985-33) should be the lower-noise device. For sub-nA noise floor requirements, the LP2985 (50 µV RMS output noise) is recommended over the AMS1117 (which is adequate for digital supplies). Total current budget at max Wi-Fi TX: ~310 mA, providing 190 mA margin against the 500 mA USB-C allowance.

## 4.2  AD5941 Analog Front-End Block

The AD5941 integrates a direct digital synthesizer (DDS) for excitation waveform generation, a potentiostat amplifier driving the CE electrode, an ultra-low-leakage trans-impedance amplifier (TIA) on the WE input, a 16-bit ADC, and an internal sequencer that can autonomously execute multi-step EIS sweep sequences without CPU intervention. This autonomous operation is critical: the RA4M1 CPU can sleep or handle USB tasks during a sweep, eliminating measurement timing jitter from software.

The 1.82V internal reference (VREF) output is decoupled with 10 µF (X5R 0603) to ground. The CE output is rated ±8 mA, sufficient for the micro-scale SPE current requirements. The internal TIA supports programmable gain resistors spanning pA to µA current ranges.

## 4.3  RA4M1 Digital Controller Block

The Renesas RA4M1 (ARM Cortex-M4F, 48 MHz, 256 KB flash) is used via the Arduino R4 Minima module for v1, allowing rapid firmware development with the Arduino ecosystem. Its RSPI peripheral drives the AD5941 at up to 16 MHz. The SCI (UART) peripheral connects to the ESP32-S3 at 921,600 baud. The RA4M1 implements COBS encoding for the UART data stream to the ESP32, providing framing without ambiguous escape sequences.

The J-Link compatible SWD debug interface (SWDIO, SWDCLK, RESET, VDD, GND) is brought out on an unpopulated 2 × 5 1.27 mm header (J_DEBUG1). A 12 MHz crystal on the R4 Minima module provides timing accuracy for SPI and UART baud rate generation.

## 4.4  ESP32-S3-MINI Connectivity Block

The ESP32-S3-MINI-1 module (15.4 × 20.5 mm) provides IEEE 802.11 b/g/n Wi-Fi and Bluetooth 5.0 (LE). The module carries its own FCC/CE/IC certification, simplifying compliance for the end product. It is placed at the board corner farthest from the analog section (bottom-right), with its on-module PCB antenna oriented toward the board edge.

GPIO0 and EN pins are connected to accessible test points for ESP-IDF firmware flashing. An unpopulated 2 × 5 1.27 mm header (J_DEBUG2) exposes TXD0, RXD0, GPIO0, EN, VDD, and GND. Bootstrapping resistors per the Espressif reference design are fitted.

## 4.5  Electrode Interface Block

Copper traces route from the AD5941 CE/WE/RE output pins directly to the PCB-integrated SPE pad zone on the right half of the board. The WE trace is the highest-impedance node on the board; it is surrounded by a guard ring trace on F.Cu connected to the AD5941 GUARD output (pin 30). This guard trace actively drives the same potential as the WE input, reducing surface leakage current on the PCB to negligible levels. No other signals are permitted within 2 mm of the WE trace without a guard barrier.

# Section 5 — PCB Layout Guidelines

## 5.1  Board Dimensions and Mechanical

| Parameter | Value |
| --- | --- |
| Board dimensions | 80.0 × 60.0 mm |
| Corner radius | R4.0 mm |
| Mounting holes | M3, Ø3.2 mm drill, 4× corners |
| Hole centers (X,Y from origin) | (6.5, 6.5), (73.5, 6.5), (6.5, 53.5), (73.5, 53.5) mm |
| USB-C connector edge location | Bottom edge, X = 31–49 mm (centered at X = 40 mm) |
| SPE electrode zone | Right half: X = 40–80 mm, Y = 6–54 mm |
| ESP32 antenna keepout | 15 × 15 mm at top-right corner, no copper all layers |

## 5.2  Zone Partitioning

- Analog Zone (top-left): AD5941, RCAL, AVDD regulation, ferrite bead. Enclosed by AGND copper guard zone.

- Digital Zone (bottom-left): RA4M1, USB-C connector, polyfuse, Schottky diode, DVDD LDO, debug headers.

- RF Zone (bottom-right corner): ESP32-S3-MINI module. 15 mm copper-free keepout in all layers under antenna.

- Electrode Zone (right half): SPE WE/CE/RE pads, guarded traces only. Absolutely no digital traces permitted.

## 5.3  Critical Trace Length Constraints

| Signal | Max Length | Impedance | Notes |
| --- | --- | --- | --- |
| WE (AIN0) | 10 mm | High-Z (fA) | Guarded trace mandatory; most critical node |
| RE (AIN1) | 15 mm | Buffered | Guard recommended |
| CE (VBIAS out) | 20 mm | Low-Z (±8 mA) | Less critical; no guard required |
| SPI (SCK/MOSI/MISO/CS) | 50 mm | 33 Ω damped | Matched ±5 mm; avoid analog zone |
| UART (RA4M1 ↔ ESP32) | 100 mm | LVTTL 3.3V | No constraints; route on B.Cu |

# Section 6 — Design for Manufacturing (DFM)

## 6.1  JLCPCB Fabrication Specifications

| Specification | JLCPCB Minimum | Recommended | Notes |
| --- | --- | --- | --- |
| Trace width | 0.10 mm | 0.15 mm (6 mil) | Use 0.15 mm default |
| Trace clearance | 0.10 mm | 0.15 mm | Increase near WE node |
| Via drill diameter | 0.2 mm | 0.3 mm | Use 0.3 mm for signal vias |
| Via annular ring | 0.13 mm | 0.15 mm | Per IPC-2221B |
| Copper to edge clearance | 0.2 mm | 0.3 mm | Critical at USB-C edge |
| Min hole-to-hole spacing | 0.5 mm | 0.8 mm | Including mounting holes |
| 0402 SMD pad size | 1.6 × 0.9 mm | Per IPC-7351 | No hand-solder pad required |

## 6.2  Board Finish and Surface Treatment

ENIG (Electroless Nickel Immersion Gold) must be specified for the entire board. This finish is mandatory for the SPE electrode pads (gold surface provides validated adhesion for nanomaterial coatings and low contact resistance) and is also beneficial for all SMD component soldering, providing long shelf life and consistent wettability. OSP or HASL are not acceptable substitutes for this design.

No-clean or aqueous wash flux processes must be specified in the assembly fabrication notes. Flux residues on or near the WE pad zone will dramatically increase surface leakage current and corrupt low-current EIS measurements.

## 6.3  Panelization for Production

For prototype runs (5–10 boards): use JLCPCB default panelization. For production volumes (> 50 boards): request V-score panelization in a 2 × 2 array (4 boards per panel, 160 × 120 mm). V-score lines at 0.4 mm offset from board edge. Add 5 mm tooling strips with fiducial markers (3 × Ø1 mm, spread across panel). Fiducials must be on both top and bottom copper layers for accurate component placement.

## 6.4  Assembly Notes

- AD5941 LGA: Reflow per Analog Devices datasheet profile. X-ray inspection mandatory on first batch of 5 boards to verify solder joint formation under the LGA.

- ESP32-S3-MINI: Follow Espressif recommended reflow profile. Module castellation pads are susceptible to bridging; use paste stencil aperture reduction of 10% on these pads.

- SPE pads: DO NOT apply solder paste to WE, CE, RE pads. These are exposed copper pads for electrochemical contact, not solder pads.

- Debug headers J_DEBUG1 and J_DEBUG2: Leave unpopulated (DNP) in production. Install only for engineering development builds.

# Section 7 — Test, Debug, and Verification

## 7.1  Mandatory Test Points

| Test Point | Location | Purpose | Acceptable Range |
| --- | --- | --- | --- |
| TP_VBUS | Near USB-C connector | Verify 5V input | 4.75 – 5.25V |
| TP_AVDD | Near AD5941 AVDD | Analog rail | 3.28 – 3.32V, < 5 mV ripple |
| TP_DVDD | Near RA4M1 VDD | Digital rail | 3.25 – 3.35V |
| TP_WE | WE trace near AD5941 | Probe WE signal | EIS-specific; scope with guard |
| TP_CE | CE output trace | Counter electrode drive | ±1.5V max during CV |
| TP_RE | RE trace near AD5941 | Reference buffer output | Stable within ±10 mV of setpoint |
| TP_AGND | AGND star point | Scope reference | 0.000V reference |
| TP_SPI_SCK | SPI SCK line | Logic analyzer hookup | 3.3V CMOS levels |
| TP_UART_TX | RA4M1 TX to ESP32 | UART sniffer | 921600 baud COBS frames |

## 7.2  RCAL Calibration Procedure

- Power on the board and allow 30 seconds for thermal stabilization.

- The RA4M1 firmware triggers the AD5941 RCAL routine automatically. Confirm via UART console: 'RCAL PASS' message with calibration coefficients.

- Verify RCAL resistor accuracy: measure TP_RCAL0 and TP_RCAL1. The voltage difference should reflect a 10 kΩ ± 10 Ω impedance.

- Firmware re-triggers RCAL automatically if internal temperature sensor reports > 5°C change from last calibration.

## 7.3  Prototype Acceptance Criteria

| Test | Pass Criterion | Tool |
| --- | --- | --- |
| Power rails within spec | All TPs within ±2% of nominal | DMM |
| AVDD noise floor | < 5 mV peak-to-peak ripple | Oscilloscope 10× probe |
| AD5941 RCAL pass | UART: 'RCAL PASS' on boot | Serial terminal |
| EIS on 10 kΩ resistor | ¦Z¦ = 10 kΩ ± 2%, phase < 1° | Firmware + Python post-proc |
| CV scan ±0.5V | Clean CV curve, no oscillation | Firmware + scope on CE |
| Wi-Fi data streaming | < 0.1% packet loss over 5 min | Laptop + Wireshark |
| BLE advertisement | Visible within 5 m, RSSI > −70 dBm | nRF Connect app |
| SPE contact integrity | < 1 Ω contact resistance on all three pads | 4-wire kelvin DMM |

# Section 8 — Signal Integrity and Noise Control

## 8.1  High-Impedance Node Design

At sub-nA current measurement levels during EIS, PCB surface leakage resistance must exceed 10 GΩ. FR4 typically exhibits 10 GΩ–100 GΩ surface resistivity in clean, dry conditions, but flux residues, humidity absorption, and ionic contamination can reduce this by several orders of magnitude. The following measures are mandatory:

- ENIG finish on all copper in the electrode zone — no OSP, no HASL.

- Guard ring surrounding WE and RE pads and their trace paths on F.Cu.

- No-clean or aqueous-wash flux process; specify in assembly notes.

- Optional: conformal coat over electrode zone (excluding WE/CE/RE pad openings) for deployments in humid environments.

- Optional for sub-pA measurements: PCB slot cut (guard groove) around WE trace to physically sever surface leakage paths.

## 8.2  ESP32 RF Noise Mitigation

ESP32-S3 Wi-Fi TX generates broadband switching noise up to 80 MHz and narrowband emissions at 2.4 GHz. During active EIS measurements, the firmware must synchronize Wi-Fi transmissions to occur only between measurement frames — never during an AD5941 acquisition window. This is implemented by having the RA4M1 assert a GPIO flag to the ESP32 during measurement (ESP32 defers Wi-Fi TX until flag is deasserted).

Physical isolation: the AVDD ferrite bead (600Ω @ 100 MHz, BLM18PG601SN1 or equivalent) attenuates any switching noise from DVDD coupling back onto AVDD. The 15 mm copper-free keepout under the ESP32 antenna prevents radiated emissions from coupling into the analog traces through parasitic board currents.

## 8.3  Guard Ring Implementation in KiCad

In KiCad PCB Editor, implement the guard ring as a copper zone polygon on F.Cu assigned to net GUARD_NET. Set the zone to flood only (no thermal reliefs). The zone must completely surround the WE and RE pads, their connecting traces, and the via stubs connecting to the AD5941 AIN0/AIN1 pins. Connect GUARD_NET net to AD5941 pin 30 (GUARD output) via a short, direct trace.

# Section 9 — Risk Register

| Risk | Likelihood | Severity | Mitigation |
| --- | --- | --- | --- |
| Analog noise from ESP32 Wi-Fi switching during EIS | MEDIUM | HIGH | Ferrite bead on AVDD; firmware-gated Wi-Fi TX; RF keepout zone |
| WE leakage current on FR4 at high humidity | LOW–MEDIUM | HIGH | Guard ring; ENIG; no-clean flux; optional conformal coat |
| AD5941 LGA solder joint defects (first batch) | MEDIUM | MEDIUM | Reflow profile per ADI datasheet; X-ray first 5 boards |
| LP2985 AVDD LDO out of stock at LCSC | LOW | MEDIUM | Qualify TLV1117 as alternative; validate noise floor with each |
| RA4M1 SPI timing incompatibility with AD5941 at 16 MHz | LOW | HIGH | Bench test prior to PCB spin; add 33 Ω series R as contingency |
| SPE nanomaterial coating adhesion failure on ENIG | LOW | HIGH | Run test coupon validation with target nanomaterial pre-production |
| USB-C current headroom insufficient under Wi-Fi TX peak | LOW | LOW | Budget 310 mA vs. 500 mA limit; 190 mA margin is adequate |
| FR4 dielectric loss degrading high-frequency EIS above 100 kHz | LOW–MEDIUM | MEDIUM | Validate at 200 kHz on first prototype; Rogers 4003C for v2 if needed |

# Section 10 — Revision Roadmap

| Version | Key Changes | Goals |
| --- | --- | --- |
| v1 (Current) | Arduino R4 Minima module; USB-C 5V; through-hole debug headers; JLCPCB 5-board prototype | Validate AD5941 EIS performance and SPE electrode geometry. Establish firmware baseline. |
| v2 | Bare RA4M1 chip; 2 MB QSPI flash; LiPo charger (BQ25180); refined SPE geometry; 70 × 50 mm board | Field-deployable unit; battery-powered operation; production BOM; reduced board area. |
| v3 | USB PD (9V→3.3V); optional calibration IC; production enclosure; CE/FCC pre-compliance | Regulatory-ready product for commercial distribution. Enhanced power efficiency. |

| NOTE | All v2 and v3 features are subject to v1 prototype validation results. Design decisions may be revised based on measured EIS noise floor, nanomaterial adhesion test results, and field trial feedback. |
| --- | --- |

# Section 11 — Document Sign-Off

This document has been prepared and reviewed by the following individuals. Signature indicates review and approval of the content within the scope described.

| Name & Role | Signature | Date | Rev |
| --- | --- | --- | --- |
| Varshini CB — Author (HW Design) |  | 08 Apr 2026 | 1.0 |
| Engineering Lead — Technical Review |  |  | 1.0 |
| Product Manager — Requirements Sign-off |  |  | 1.0 |
| Manufacturing Lead — DFM Approval |  |  | 1.0 |

End of Document — VL-HW-001 Rev 1.0

© 2026 VidyuthLabs Pvt. Ltd. All Rights Reserved.